## Design considerations
Open Source Ventilator Project Pakistan Specifications
S.No | Design Parameter | Design Range | S.No | Patient Safety, monitoring and alarms
1 | Pressure Inspiratory | up to 35 cmH2O Default: 30 cmH2O | 1 | Alarm – Minute Ventilation | Yes, Low/High | Video to understand ventilators
2 | Pressure Expiratory (PEEP) | 5 to 25 cmH2O in increments of 5 | 2 | Alarm – Peak Pressure | Yes | https://youtu.be/7Y7JfeQMz_s
3 | Respiratory Rate | 10 to 30 bpm Increments of 2 (only in mandatory mode) | 3 | Alarm – Low expiratory pressure | Yes
4 | I:E Breathing Ratio, Inspiratory:Expiratory | 1:1 to 1:2 & 1:4 | 4 | Alarm – O2 Disconnection | Yes
5 | FiO2 (Inhaled O2 %) | Minimum 50% to 100% Preferred 20% to 100% in 10% steps | 5 | Alarm – PEEP pressure (disconnection) | Yes | Primary Sources
6 | Ventilation Triggering | Timed or Flow or pressure | 7 | Alarm – Manual off when in mandatory mode | yes
7 | Tidal Volume (Vt) | Must have at least one setting of 400ml +/- 10 ml Range from 100-1000ml | 8 | Alarm – Tidal Volume | Yes, Low/High | Uk govt. Specifications
8 | Tube/Patient Connectivity | Standard connections | 9 | Monitoring – Tidal Flow | Yes
9 | O2 Connectivity | Standard O2 connections | 10 | Monitoring – Respiratory Rate | Yes | Open Lung Project
10 | Accuracy | <10% for volume and pressure, 1 breath/min | 11 | Monitoring – PEEP | Yes
11 | Primary Power | 110v and 220v AC (60Hz!!) | 12 | Monitoring – Plateau Pressure | Yes | Dr. Ali Shoaib | Liaqat national
12 | Aux Power | Battery | 13 | Monitoring – O2 Concentration (FiO2) | Yes
THIS SHEET IS FOR DOCTORS ONLY.
Step by Step Procedure for a Ventilator of above design specifications
| Dr. Mehreen malik. | Anesthetist and critical care specialist at Aku:
set tidal volume 5-8 ml per kg
set respiratory rate 15 breaths per minute
set peep starting at 5 initially can go upto 30
set I:E ratio 1:2 initially
set max pressure limit of 30 cm h2o
these are the simplest parameters that can be set initially when attached to the ventilator
this can be attached to artificial lung to test.
the process after that is when ventilator will deliver first tidal volume breath to the patient and give time for chest to recoil.
after recoil of chest the ventilator will again give another breath. so we need to just decide how much volume of air is delivered and set respiratory rate per minute. taking care of pressure that doesn't go beyond point to cause lung injury.
it's like a simple Manley ventilator.
| Dr. Akbar | ER physician.
The patient is intubated and adequately sedated and relaxed.
Perform Short Self test to check ventilator is working or not
Select `New patient` and enter patient weight which should be by formula ♂: PBW (kg) = 50 + 2.3 (height (in) – 60) ┃ ♀: PBW (kg) = 45.5 + 2.3 (height (in) – 60)
Select Assist control mode
Select Volume Control (Allows you to calibrate Tidal Volume, PEEP, Resp. rate and FiO2), FiO2 should remain 100% to begin
Set Tidal Volume 4-6 ml/kg
Select desired respiratory rate as per clincal scenario. Normally set to 12-20 breath/min.
Caculate desired Minute Ventilation (TV x RR) and keep it approximately 6-7L/min
Select PEEP to 5cm H2O. For ARDS PEEP may need to be set higher to 7-10 cmH2O but very carefully as it can cause Alveolar recruitment and/or trauma.
Select Flow (Volume) trigger to 2.0 or 3.0
Set Inspiratory flow to 60L/min
Select waveform to Square or Ramp shape
Check to see I:E ratio is calibrated to 1:2
Mechanical ventilation is a dynamic process so patient is continually reassessed and if required Arterial Blood Gases are performed for guidance.
In ARDS, as well as other patients, maintaining a Pplat <30cmH2O is key which is found out after doing an inspiratory hold
https://www.youtube.com/watch?v=seU-upuP2wY
If you were to select Pressure Control after selecting Assist Control mode (Allows you to calibrate Insp. pressure, Resp. rate, PEEP, FiO2) - FiO2 should remain 100% to begin.
Select Inspiratory Pressure (Pi ) to 15 - 20 cm H2O
Select i Time to 0.7 - 1.0 second
Select respiratory rate 12 - 20 breath/min.
Select PEEP to 5 cm H2O
Select Flow (Volume) trigger to 2.0 or 3.0
the tidal volume, peak Insp. pressure and Inspiratory/expiratory pressure will dynamically change on screen and you may recalibrate as per desired values according to clinical assessment and/or arterial blood gas values.
https://www.youtube.com/watch?v=ainxKVsVs1s
| Ihtesham Ul Haq | ICU Technologist
Step 1. Calibration of Ventilator with special conective tube of both values
Step 2 . Sedation and relaxation of Patient
Step 3. Prone Postion of ARDS patient ( 7 to 8 hours ) then it can be change
Step 4. (i)Intial Tidal volume 600 ml
ii)peep 5cmH2O or above depend on patient condition
iii)Fio2 will be given 100% intial to Covid19 and to ARDS 40% depends on patient condition
References https://youtu.be/tarWCSbdcGA
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4828494/
iv)flow 60l/min
v)Respiratory rate 15bpm
vi)Mode PSV
v)I.E ratio 1:2
vi)Inpiratory Pressure 15-20 cmh2o
Dr moazzam resident anesthesia , sicu and pain management
enter basic mode of ventilator like CMV,SIMV,spontaneous
enter peep 0-15 cm of water
set I:E ratio 2:1,3:1,1:1,1:2,1:3
other parameters same as described above
CMV: controlled mechanical ventilator
SIMV: synchronized intermediate mechanical ventilator
spontaneous:patient breath spontaneously
Note on G5: Failsafe valve needed at 40cmH2O -Ahsen Shafique
Note on D15: Isnt 220Vrms, 50Hz used in Pakistan? -Ahsen Shafique
Note on D20: We first set tidal volume 500 ml of a 70 kg man. Start from 5 ml per kg. Set respiratory rate 15 breath per minute standard Set I:E ratio of 1:2 Set peep of 5 These are basics -Mehreen Malik Attach the ventilator to the patient and start ventilation -Mehreen Malik
Note on C21: Can Doctors please also give ventilator guidelines and settings specifically for ARDS? The ventilator specs must be able to respect these as Covid patients as likely to have ARDS. For instance UK guidelines recommend Pressure control, low tidal volumes and very high PEEP pressures. -Adil Malik